设备技术规格书

# Equipment Technical Specification

设备名称：库仑测厚仪

Equipment name: Plating thickness tester

型号/Model：Galvanotest 3000

厂家及产地/ Manufacturer and Production Place：德国EPK/ ElektroPhysik

主要技术参数/Main technical requirements：

- 测量范围/Measure range： 0.05μｍ到75μｍ

- 准确测量电镀件（铜、镍、铬等）各层厚度/Accurately  measure the plating thickness(Cu/Ni/Cr);

- 测量面积/Measuring  area:	B:	4mm2 (直径2.26mm)     C:	1mm2 (直径1.13mm)  	          								D:	0.25mm2 (直径0.56mm)	K:	电解杯(为测量小部件和导线用)

- 测量不确定性 /Measuring uncertainty:

4mm2测量面积：((％读值)

1mm2测量面积：((％读值)

0.25mm2测量面积：((％读值)

如果没有标准板，可认为UST为10％。

*UST－镀层厚度标准板的容差(百分数表示)。大数情况下为±5％。

| Rubber gasket / mask | Diameter | Measuring area | Measuring uncertainty* in percent of reading |
| --- | --- | --- | --- |
| Rubber gasket B | 2,26 mm | 4mm2 | 6% |
| Mask C | 1,13 mm | 1mm2 | 9% |
| Mask D | 0,56 mm | 0.25mm2 | 20% |

* The specifications are based on calibration of GalvanoTest 3000 using ElektroPhysik thickness standards.

- 其他参数

可测量超过70种涂层/基体组合，请参阅电解液选择表

Measurable coating/base combinations more than 70, please refer to the electrolyte selection table

溶解速度/De-plating speed：0.3～40(m/min

单位/Measuring unit：英制或公制(inches／μm)

厚度显示/ Display of readings：4位浮点读值/4-digits digital display with floating point 。

16位字母数字显示/ Alphanumerical display 16-digits ：

显示仪器调整值，如金属，溶解速度等/display of instrument settings and operating steps。

指针表：监视溶解过程Analogue instrument for monitoring the de-plating process 。

数据存储/ Data memory：

18种不同金属选择，可存储2000个读值/max. 2,000 readings storable in a maximum of 18 application memories。

操作：按键/Operation via key pad

备用接口/Data ports ：RS-232

电源Power supply：110V／220V±10%  50/60Hz /10W

重量/Weight：仪器/ instrument：  2.5Kg	   测量支架/ precision holder： 2.5Kg

备件和消耗材料/Spare parts & expendables

- 测量电镀件（铜、镍、铬等）相关电解液及电解头/Relative car plating thickness measure electrolyzing solution and measure head;

- 标准配置/Standard supply

Galvanotest3000仪带模拟表、RS232C接口、内置循环泵的测量支架

-GalvanoTest 3000 processing unit

-interface for MiniPrint 4100 data printer or PC

-measuring stand 3000 with integrated circulation pump

I型含内部喷嘴的电解池/outer measuring cell and inner measuring cell, type I

2个密封垫 B型(4mm2) /2 rubber gaskets B (4,00 mm2)

2个用于C、D片的密封垫,直径1.5mm  /2 rubber gaskets (.1,5mm) for sealing the masks

20片测量面封贴片/20 masks  C (1,00 mm2)

20片测量面封贴片/20 masks D (0,25 mm2)

橡皮铅笔/rubber pencil

5 x100 ml可按客户要求的电解液/5 x100 ml electrolyte solution, a sorted to customer requirements

1个100 ml装使用过的废电解液的瓶子/1 bottle for used electrolyte solution, 100 ml

1个100 ml用于清洗电解池的喷雾瓶  /1 washing bottle, 100 ml

吸水纸1包/1 package absorbant paper

操作说明书1份 /operating instructions including works certificate

- 推荐备件/Recommended accessories

使用不同电解液时的I型或II型内部喷嘴/inner measuring cell type I or II, for working with a second electrolyte solution

定位臂/ positioning arm

替换泵管/ spare pump tube

带连接电缆的电解液杯/ electrolyte cup incl. connecting cable

导线镀层测量夹/wire holder for measuring wires

小部件夹具/ small parts holder

7瓶100ml电解液的瓶架/ bottle stand for holding up to 7 electrolyte solution bottles, 100 ml each

校准板/coating thickness standards

用于5个内部喷嘴的测量支架/ measuring cell stand for holding up to 5 inner measuring cells

替换密封垫/spare rubber gaskets    B (4 mm2)

封贴片/masks  C (1 mm2)

封贴片/masks  D (0.25 mm2)

定制密封垫/modified rubber gasket D/Au (0,25 mm2)

备用吸管/spare pipette

附加电解液/electrolyte selection

备用橡皮铅笔/spare rubber pencil

定心器/centering device (Z2)

打印机MINIPRINT 4000/ external accu battery powered data printer MiniPrint 4100

打印机MINIPRINT4000的连接电缆/ connecting cable for MiniPrint 4100 data printer

连接记录仪的电缆/connecting cable for writer

连接PC或兼容打印机的RS232电缆/ RS 232 connecting cable for PC or printer

MSoft7000数据传输程序/MSoft7000 data transfer program

附录：电解液(E)和镀层厚度标准(ST)选择/Elektrolyt selection table

| Coating 底材 镀层 Substrate | Ag | Au | Cd | Cr | Cu | Brass | Ni | 化学镍Current-less Ni | Pb | Pb/Sn | Sn | Sn/Zn | Zn |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | 银 | 金 | 镉 | 铬 | 铜 | 黄铜 | 镍 |  | 铅 | 铅/锡 | 锡 | 锡/锌 | 锌 |
| 非金属 Non metal | E4 |  | E5 | E11 | E4 | E4 | E14 | E17 | E15 | E4 | E7 | E7 | E6 | * |
|  |  |  | 1 | 8 |  |  |  |  |  |  |  |  |  |  |  |  |  | 8 |  | 8 |  |  |  |  |  |  | ** |
| 银 Ag |  |  |  |  |  |  |  |  | E15 |  |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 8 |  |  |  |  |  |  |  |  |
| 铝 Al | E4 |  | E5 | E11 | E4 | E4 | E14 | E17 | E15 | E4 | E11 |  | E6 |
|  | 8 |  |  |  | 8 |  | 8 |  | 8 |  | 8 |  | 8 |  |  | 8 |  | 8 |  | 8 |  | 8 |  |  |  | 8 |
| 青铜 Bronze | E8 |  |  |  |  |  |  |  |  |  | E7 |  |  |
|  | 8 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 8 |  |  |  |  |
| 镉 Cd |  |  |  |  |  |  |  |  |  |  | E7 |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 8 |  |  |  |  |
| 铜 Cu | E8 |  | E5 | E7 |  |  | E14 |  | E15 | E4 | E9 | E7 | E6 |
|  |  | ST | 1 | 8 |  | ST |  | ST |  |  |  |  |  |  | ST |  |  |  |  | 8 |  | 5 | ST | 8 |  | ST |
| 钢铁 Steel | E4 |  | E5 | E11 | E4 | E4 | E14 | E17 | E15 | E4 | E7 | E7 | E20 |
|  |  | ST | 8 | 8 |  | ST |  | ST |  | ST |  |  |  |  | ST |  |  | 8 |  | 8 |  |  | ST | 8 |  | ST |
| 黄铜 Brass | E8 |  | E5 | E7 | E12 |  | E14 |  | E15 | E4 | E9 | E7 | E6 |
|  |  | ST | 1 | 8 |  | ST |  | ST | 8 |  |  |  |  |  | ST |  |  | 8 |  | 8 |  | 5 | ST | 8 | 5 | ST |
| 镍 Ni | E8 |  | E5 | E11 | E4 | E4 |  |  | E15 |  | E7 |  | E6 |
|  | 8 |  | 1 | 8 | 8 |  |  | ST | 8 |  | 8 |  |  |  |  |  |  | 8 |  |  |  | 8 |  |  |  | 8 |
| 硅 Si |  |  |  |  |  |  | E14 |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  | 8 |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 锌 Zn |  |  |  |  | E12 |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  | 2 | ST |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 柯伐合金 Kovar |  |  |  |  | E4 |  |  |  | E15 |  |  |  |  |
|  |  |  |  |  |  |  |  |  | 8 |  |  |  |  |  |  |  |  | 8 |  |  |  |  |  |  |  |  |
| 镍-银 Ni-Ag | E8 |  |  | E7 |  |  |  |  |  |  | E7 |  |  |
|  | 8 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 8 |  |  |  |  |
| 铸锌 Cast Zn |  |  |  |  | E12 |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  | 8 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

- 金镀层应客户要求而定 。

For electrolyte selection and instrument adjustment lease refer to our customer support.

- 见说明书9.8节(Cu／Zn) / See section 8.9 “Copper on Zinc” 。

- 见说明书9.9(锡／黄铜) / See section 8.10 “Tin on Brass” 。

- 适用于锡／锌，78％/22%  / For Tin/Zinc 78%/22% 。

- 这种组合用II型电解池 / Please use the inner measuring cell type II 。

(8)	 这种基底镀层组合需特别仔细，请与本公司联系。

Requires carefull measurement. Please contact our customer support if necessary .

(ST) 可供镀层厚度标准 /Coating thickness standards available on stock 。